## Slide 1
In-Situ Measurement and Detection of Diffusible Hydrogen in Ferritic and Austenitic Stainless Steels
T. Hillman1 , M. Leites1 , B. Clowes2 , R. Bernard2 , D.L. Engelberg1
1Department of Materials, School of Natural Sciences, Metallurgy & Corrosion, The University of Manchester, Oxford Road, M13 9PL, UK
2Sellafield Ltd., Sellafield, Seascale, CA20 1PG, United Kingdom

## Slide 2
Contents
Overview
  Hydrogen (forms and conditions)
  Hydrogen Detection Routines
  Laser-Induced Breakdown Spectroscopy
LIBS Systems
Cumulative Shot Behaviour
Handheld Calibrations
LIBS-Devanathan-Stachurski
Conclusions

## Slide 3
Hydrogen – Introduction (Forms & Conditions)
Hydrogen Gas (H2): molecule size / gaseous size (0.24 nm)
Permeation/diffusion in polymers.
Hydrogen Atoms (H): dissolved in metal (atom size: 0.1 nm)
Mobility: in metal octahedral/tetrahedral sites for lattice diffusion; dependent on trap sites & their availability.
Hydrogen in ionic form:
(H+) - Acidity in water (hydronium ion / controls pH-value); pitting & crevice corrosion (anoxic conditions).
(H+) - ion-conductive membranes (Nafion – Sulfonated tetrafluoro ethylene; but also H2)
(H-, H-σ) - Metal-hydrides NiH, ZrH, MgH2 etc. (0.28 nm)
Hydrogen bound in compounds (covalent):
Methanol (CH3-OH); alcohols, organics other petrochemical / polymers.
NH3, other feedstock/storage/transport media.
Natural Gas H2S (odorant) / Methane (CH4), Ethane (C2H8),…
H./H2 generation at cathodic sites

## Slide 4
Overview Ex/In-Situ Techniques – H Concentration / Location Measurement
Lab Techniques
Atom Probe Tomography
SIMS/Nano-SIMS
Scanning Kelvin Probe
GD-OES
Destructive – Extraction
Hot gas & melt extraction
TDS/TDA
Microstructure
Microprinting (Ag decoration)
PANI
Electrochemistry
Devanathan-Stachurski (H-Permeation)
“Quasi-non-destructive”
Laser-Induced Breakdown Spectroscopy
nanosims.asiaa.sinica.edu.tw/
Horiba Scientific
M. Koyoma, et al, Scripta Materilia, 129 (1) (2017) 48-51

## Slide 5
Laser-Induced Breakdown Spectroscopy (LIBS)
LIBS
Nd:YAG laser (1064 nm)
Argon shielding gas
Plasma
plume
Excitation
Emission
Relaxation
Detection
Spectrum (Ex. 316L)
Hα

## Slide 6
Doppler broadening (Gaussian):
Stark broadening (Lorentz):
1Reinhard Noll, Laser-Induced Breakdown Spectroscopy Fundamentals and Applications. Springer Heidelberg Dordrecht London New York, 2012.
1
LIBS – Raw Data Processing
Where necessary baseline correction
Normalisation of spectra [0;1] (if required)
Peak fitting with combined Voigt function
Analyse H at Hα peak at 656.25nm

## Slide 7
Manchester LIBS Systems

## Slide 8
Z300 Handheld Calibration
Calibration for in-situ H-LIBS use
SciAps Z300
Hydrogen stainless steel standards (6.23ppm – LECO; 2.3ppm, 0.95ppm – SciLab)
Integration time, delay = 1 ms, 250ns
Resolution < 0.1mm x 1 mm
Ar purge
Normalize Hα at 656.2nm to Ar(I) at 763.5nm
Laser power = 5-6 mJ pulse-1
Chart: [H]/ppm
| 0.95 | 2.2999999999999998 | 6.23
Normalised Intensity | 0.18631 | 0.37382 | 0.75725

## Slide 9
+
+
+
+
+
+
–
–
–
–
–
–
Hydrogen Permeation-type LIBS
Thermometer
REFLUX
Pt counter
Heating
+
–
Sample
0.5M Na2SO4
LIBS acquisition

## Slide 10
Reflux condenser
Thermometer
SciAps Z300 handheld LIBS
1064 Nd-YAG laser
5-6mJ pulse-1
1 shot per location
2 locations
Heating (80°C)
Hydrogen Permeation-type LIBS

## Slide 11
Hydrogen Permeation-type LIBS
Intensity = 0.1058[H] + 0.1049
??
Normalise to Ar(I) at 763.5nm
??

## Slide 12
Cumulative Shot Hydrogen Analysis
To investigate the efficacy of multi-shot/depth profiling LIBS for hydrogen analysis
LIBScan 100 Ultra
Hydrogen stainless steel calibration standards (6.23ppm – LECO DH603; 2.3ppm, 0.95ppm – SciLab Eltra OH)
30 consecutive shots per sample
Resolution < 0.5mm
140mJ, 1Hz laser pulse
Hα peak at 656.25nm normalised to Fe(I) line at 495.75nm

## Slide 13
Cumulative Shot Hydrogen Analysis
High initial normalised intensities
Large variance in errors
  Increased scattering of incident and re-emitted light?

## Slide 14
Cumulative Shot Hydrogen Analysis
High initial normalised intensities
Large variance in errors
  Increased scattering of incident and re-emitted light?
General decrease in intensity
  H depletion and enrichment (at greater depths?)

## Slide 15
Conclusions
Hydrogen detection with LIBS still being refined
    Rapid (<5s), local (<<0.5mm) and in-situ H detection possible. Normalisation techniques still being explored…
Standardisation possible, yet still room to improve (global standardisation of a local technique)
Cumulative shots reveal some “movement” of H atoms within steel standard matrix
Devanathan-type LIBS possible – cannot expect a classical permeation curve just yet!
NEXT STEPS:

## Slide 16
Thank you for listening

## Slide 17
Cumulative Shot Hydrogen Analysis